# Corial ICP Operation Instructions

# At

# University of Notre Dame

# Department of Electrical Engineering

Contacts

For problems, clarification of procedures, or general information pertaining to this machine please contact send email to   staff@nd.edu

In case of emergency, please contact the Notre Dame Security at

911

General Precautions

- Serious Equipment Damage will occur if you do not use the bubble level to verify correct puck placement inside the load lock

- Serious Equipment Damage will occur if you leave the bubble level inside the load lock

- Pucks can become very hot and are very fragile – please do not touch them when hot

- The Five (5) piece quartz puck MUST have a full 100 mm wafer in it in order to use.  BS Helium will not function without a full wafer.

- Pucks (Quartz and 5-piece wafer) are never to be left on tables or on top of the machine.   They must remain inside the blue boxes or inside the machine.  Five pieces are 2 – quartz rings, 1 blue oring, 2 metal plates

- You must include a 5 second bias plasma ignition step in order to assist with ICP plasma ignition.  You must also include two post etch pumps and evacuation steps

- You MUST return cathode Temp to “Set Standby” temperature when finished.

- ICP bias voltage CAN NOT Exceed 2000V – errors will occur

- You MUST be present at the machine and monitoring the system anytime RF is turned ON

## Running a Sample

## Enable the machine in iLab

- On the ICP computer, go to “Handler”, then “Transfer”, and “Vent Lock” to vent the Load Lock chamber.

- Once at ATM – gently open the lid

- Select either:

  - Single piece quartz puck

  - Five (5) piece 100 mm wafer puck and assemble correctly – two quartz pieces, 1 blue oring and 2 metal pieces compose up the 5-piece puck.  It Must have a full 100 mm wafer in it in order to use.  BS Helium faults will occur if there is not a wafer.

- Install puck of choice in the load lock and verify rotation.

- SET THE BUBBLE LEVEL ON AND VERIFY 50% or more of the bubble is in the middle of the center

- REMOVE BUBBLE LEVEL – failure to not remove this can cause serious equipment damage.

- Close Lid

- On the computer go to “Handler”, then “Transfer” and “Pump Lock”

- Allow load lock pressure to reach “Pumping Idle” on status screen – this may take about 1 minute

- Select “Transfer” and allow the puck to transfer into the system.

- Go to the “Process” tab – then “Start Job”

- Load Recipe and enter Job ID name

- Select “Start Job” button

- Allow process to complete

- From “Handler” – “Transfer” tab – select “Transfer” to return the puck back to the load lock.

- Allow the load lock to return to “Pumping Idle” status

- From the “Handler” “Transfer” tab – select “Vent Lock”

- Allow the load lock to vent to atmosphere.

- Open lid – VERIFY PUCK IS NOT TOO HOT TO HANDLE and remove only when safe to do so.

- Place puck into blue box

- Close lid

- Select “Pump Lock” when finished

- Set “Standby Temperatures” when finished

- Clean up the equipment, samples, trash and accessories you may have used.

- Disable machine in iLab.

Making a Process Recipe

- From the “Recipe” mode, go to “Edit Recipe”, then “Load”

- Select 1_Lab Oxygen Clean

- Change Recipe Name and Select “SAVE” after confirming “Engineering” type

- Place net ID in Description section – additional info (ER, DC Bias) is optional

- To modify recipe – you will need to modify the following steps:
RF On and Main p and main etch steps to the gas flow, pressure and RF Powers that you desire.

  - Initial – Temperature Definition Step

  - RF On – Gas Flow, Pressure and RF Bias Strike power

  - Etch – Gas Flow, Pressure, RF Bias, RF ICP power, etch time