# Heidelberg MLA150 Operating Instructions

## General Precautions

Contacts

For problems, clarifications of procedures, or general information pertaining to this machine, please contact the lab staff:

Lab staff email:  	staff@nd.edu

In case of emergency, Please contact the Notre Dame Security  at 1-5555 or 911

# Warnings:

# 1 – Do not leave alignment jigs inside the machine – serious equipment damage will occur if left on stage.

# 2 – Do not use pneumatic alignment with small pieces

# 3 – min sample size is 10 mm x 10 mm

# 4 – No Metal Tweezers

System will follow the workflow:

## Load Substrate

## Enable the machine in iLab

## Ensure stage is at unload position

## Open door

## Insert jig and align for 4” wafer

## Insert Flat on alignment jig

- Apply vacuum

- REMOVE JIG!!!!!

- Close door

- Press Load substrate

- Continue

  - System moves to load under center

- Confirm substrate is centered under write head

- Continue

  - System will focus

  - System will center substrate

- Alignment if second layer – See alignment section

- Start Exposure

## Converting a Design

- Save File to desktop/design/file type/create user folder

- Open GUI HIMT Converter (X) or Convert Design in JOB

- File

- New Job

- Give it a name – nothing fancy, no characters

- Add – select file type (ex:  gdsii)

- Open folder

- File – Open

- STEPPER MASKS – must increase scale (X and Y) to 5 to print 5x

- Select layer to convert

- Create Default

- Viewer – allows you to see conversion file

- File – Save Job as HIMT.job file for use in exposure (home/convert/gui/jobs/yourfile.job)

- Complete Task (exit)

## Setup Job

- Select new or load an existing job.

- All Jobs  – fill in Green information

  - You do NOT have to select a resist

  - You can change dose prior to writing

  - If resist and laser selected do not match you will get a warning but the system will write with the laser that has been selected.

- Types of jobs:

  - Standard – What is written in gdsii or conversion file

  - Series – Focus and/or Exposure matrix

- Select substrate type

- 1st layer does not need Alignment Settings

- Subsequent layers will require Alignment Settings

- BLUE line  exposure layer

Alignment

## Load job and sample

## Move to 1st alignment markter (cross or square)

- Measure

- Accept Position if good

- System moves to 2nd alignment marker

- Measure

- Accept position if 2nd alignment is good

Small Parts and Pieces

## Place small parts jig in machine

## Install sample – push towards left side of alignment jig

## Open job

## Select “Small” for substrate

## Select exposure – change alignment mode to MANUAL if second layer

## Select Optical Focus  Pneumatic Focus WILL NOT WORK on small pieces

## Load sample

## Might have to change magnification on microscope

## Center on piece

## Running a Focus Exposure Matrix

- Load Job – MLA_1000_L1 or FE2MIN

- Configure Substrate for correct information

- Select SERIES

- Select Defocus and/or Exposure – select range and increments

## Mask Making Process

- Include substrate cost in iLab reservation

  - Listed under “add additional service charges”

- REMEMBER – 5x scaling is needed (X and Y) in job conversion if you are making a stepper mask!!

- Place 0, 0 in the center of your design

- All Masks will be inverted – use caution with text unless you flip horizontally

- Expose 4” or 5” mask substrate in MLA system

- Develop – AZ917 – 25 to 30 sec

- Rinse in DI – Inspect

- Chrome Etch ~65 sec CR-4 (loading dependent)

- Rinse in DI – inspect with transmitted light

- PR removal – rinse in ACE/IPA

- Drytek 5 minutes

Alignment Markers

Ex below – From Center of Die:

- Top Cross at (0, 5000)

- Bottom Cross at (0, -5000)

Standard Cross is 16 um wide at center

Square – Info needed